## Precision & MU
n | 29
k | 2 | Expansion Coefficient
Reproducibility sd | 0.12525 | Reproducibility from ANOVA
Mean | 3.4778 | Average of samples
Precision | 0.25051 | =k*Reproducibility
Bias | -0.32469 | Insignificant | Average of Bias*mean of sample/100
UoBias | 0.29587 |  | Uncertainty*mean/100 | 3.478 | ug/100 g
|  |  |  | 34.78 | ug/kg
MU ± | 0.64259 | ug/100 g |  | 0.03478 | mg/kg
| 18.48% | Howitz MU | 3.478±1.8 | 53.05%
Replicates | 0.35427
Rounding | 0.01
| 3 | 0.371
| 2*U(Bias) < abs(aveBias) | significant
| 2*U(Bias) > abs(aveBias) | insignificant

## Repeatability
|  |  |  | n | 2
|  |  |  | k | 4.3027
|  |  |  | Repeatability sd | 0
|  |  |  | Mean | #DIV/0!
|  |  |  | Duplicates | 0
|  |  |  | Rounding | #NUM!
Matrix | Ref | REPORTED_UNITS | A | B | Diff | sqr | Omit | Outlier
[5 empty] | 0 | 0 | No | No
[5 empty] | 0 | 0 | No | Yes

## ANOVA
2015-02-24 | 2015-02-25 | 2015-02-26 | 2015-02-27 | 2015-03-26 | 2015-03-27 | 2015-03-28 | 2015-03-29 | 2015-03-30 | 2015-03-31 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
3.4574 | 3.4019 | 3.6337 | 3.4225 | 3.3651 | 3.3587 | 3.5756 | 3.4909 | 3.6915 |  | 3.4574 | 3.4019 | 3.6337 | 3.4225 | 3.3651 | 3.3587 | 3.5756 | 3.4909 | 3.6915
3.3826 | 3.4582 | 3.5829 | 3.4631 | 3.2554 | 3.3554 | 3.4919 | 3.5154 | 3.5716 | 3.5688 | 3.3826 | 3.4582 | 3.5829 | 3.4631 | 3.2554 | 3.3554 | 3.4919 | 3.5154 | 3.5716 | 3.5688
3.4542 | 3.3728 | 3.4949 | 3.4442 | 3.2254 | 3.3977 | 3.4335 | 3.5874 | 3.7729 | 3.6304 | 3.4542 | 3.3728 | 3.4949 | 3.4442 | 3.2254 | 3.3977 | 3.4335 | 3.5874 | 3.7729 | 3.6304
3.5022 | 3.5022 | 3.5022 | 3.5022 | 3.5022 | 3.5022 | 3.5022 | 3.5022 | 3.5022 | 3.5022
Anova: Single Factor
SUMMARY
Groups | Count | Sum | Average | Variance
2015-02-24 | 3 | 10.2941 | 3.4314 | 0.0017892
2015-02-25 | 3 | 10.2329 | 3.411 | 0.0018863
2015-02-26 | 3 | 10.7115 | 3.5705 | 0.0049291
2015-02-27 | 3 | 10.3298 | 3.4433 | 0.00041302
2015-03-26 | 3 | 9.8459 | 3.282 | 0.0054048
2015-03-27 | 3 | 10.1118 | 3.3706 | 0.000555
2015-03-28 | 3 | 10.501 | 3.5003 | 0.0051007
2015-03-29 | 3 | 10.5937 | 3.5312 | 0.0025162
2015-03-30 | 3 | 11.036 | 3.6787 | 0.010251
2015-03-31 | 2 | 7.1992 | 3.5996 | 0.0019013
|  | Average | 3.4818
ANOVA
Source of Variation | SS | df | MS | F | P-value | F crit
Between Groups | 0.35956 | 9 | 0.039951 | 11.2301 | 6.8141e-06 | 2.4227
Within Groups | 0.067592 | 19 | 0.0035575
Total | 0.42715 | 28
Standard Deviation
Repeatability sd |  | Sr | 0.059644
|  | Sb | 0.11014
Reproducibility sd |  | Si | 0.12525
Intermediate precision |  |  | 0.25051
7.195%
F < F critical | homogeneity and reproducibility
P-value > 0.05 | greater 95% confidence

## Bias
| U(Ref Conc), % | 7.8865
| Average Bias, % | -9.3361
| sd (Bias), % | 3.1906
| U(Bias), % | 8.5075
| Bias Significance | Insignificant
| 2*U(Bias) < abs(average Bias) | significant
| 2*U(Bias) > abs(average Bias) | insignificant
Type | Reference | Date | Unit | Matrix | Reference Mean | sd | n | Lab Result | % sd | Bias | % Bias
SRM | NIST 1849a |  | ug/100g | IF | 4.82 | 0.85 | 5 | 4.49 | 17.6349 | -0.33 | -6.8465
SRM | NIST 1849a |  | ug/100g | IF | 4.82 | 0.85 | 5 | 4.5 | 17.6349 | -0.32 | -6.639
SRM | NIST 1849a |  | ug/100g | IF | 4.82 | 0.85 | 5 | 4.27 | 17.6349 | -0.55 | -11.4108
SRM | NIST 1849a |  | ug/100g | IF | 4.82 | 0.85 | 5 | 4.15 | 17.6349 | -0.67 | -13.9004
SRM | NIST 1849a |  | ug/100g | IF | 4.82 | 0.85 | 5 | 4.44 | 17.6349 | -0.38 | -7.8838
4.37
0.15379
3.519%

## Linearity
Conc | Response
0 | 0
1 | 1.02
2 | 2.03
4 | 4.1
6 | 6.1
10 | 10.05

## LOD-LOR
sd | 0.0040222 | #DIV/0! | #DIV/0! | #DIV/0!
LOD | 0.012066 | #DIV/0! | #DIV/0! | #DIV/0!
LOR | 0.040222 | #DIV/0! | #DIV/0! | #DIV/0!
Result | 1 | 2 | 3 | 4
1 | 0.01
2 | 0.02
3 | 0.01
4 | 0.015
5 | 0.014
6 | 0.01
7 | 0.011
8 | 0.02
9 | 0.014
10 | 0.018

## Raw Data
2015-02-24 | 2015-02-25 | 2015-02-26 | 2015-02-27 | 2015-03-26 | 2015-03-27 | 2015-03-28 | 2015-03-29 | 2015-03-30 | 2015-03-31
3.4574 | 3.4019 | 3.6337 | 3.4225 | 3.3651 | 3.3587 | 3.5756 | 3.4909 | 3.6915
3.3826 | 3.4582 | 3.5829 | 3.4631 | 3.2554 | 3.3554 | 3.4919 | 3.5154 | 3.5716 | 3.5688
3.4542 | 3.3728 | 3.4949 | 3.4442 | 3.2254 | 3.3977 | 3.4335 | 3.5874 | 3.7729 | 3.6304
3.5022 | 3.5022 | 3.5022 | 3.5022 | 3.5022 | 3.5022 | 3.5022 | 3.5022 | 3.5022 | 3.5022

## Latest_1922016
IRM001A-1
Date | Bound Inositol Result(mg/100g) | Free Inositol Result(mg/100g) | Total Inositol Result(mg/100g)
2015-12-02 | 10.4913 | 33.39 | 43.8813
| 9.1591 | 38.11 | 47.2691
9.4486
| 9.3861 |  |  |  | IRM001A-1
2015-12-08 | 8.3694 | 31.17 | 39.5394
| 7.8145 | 34.53 | 42.3445 | 2015-12-02 | 2015-12-08 | 2015-12-21 | 2015-12-22 | 2016-01-12 | 2016-01-18 | 2016-01-20 | 2016-02-02 | 2016-02-09 | 2016-02-16
| 8.4619 | 30.16 | 38.6219 | 43.8813 | 39.5394 | 36.2376 | 42.5642 | 42.8258 | 43.4899 | 42.2452 | 41.6217 | 41.2357 | 39.1827
2015-12-21 | 7.5776 | 28.66 | 36.2376 | 47.2691 | 42.3445 | 39.5779 | 42.8939 | 43.1453 | 43.9376 | 44.6968 | 43.0943 | 42.9248 | 39.7552
| 8.0179 | 31.56 | 39.5779 |  | 38.6219 | 39.6368 | 42.9764 | 42.5869 | 44.5447 | 41.7684 | 42.3114 | 41.6344 | 39.1991
| 7.4968 | 32.14 | 39.6368 |  |  | 38.1873 | 42.1186 | 42.9487 | 44.2861 | 40.9185 | 42.4406 | 44.0386 | 38.7817
| 6.4373 | 31.75 | 38.1873
2015-12-22 | 11.4042 | 31.16 | 42.5642
| 11.7239 | 31.17 | 42.8939
| 11.1964 | 31.78 | 42.9764
| 11.4986 | 30.62 | 42.1186 | 2015-12-08: Anova: Single Factor
2016-01-12 | 8.5758 | 34.25 | 42.8258
| 8.8153 | 34.33 | 43.1453 | 2015-12-08: SUMMARY; 2016-01-12: mg/100 g
| 8.5969 | 33.99 | 42.5869 | 2015-12-08: Groups; 2015-12-21: Count; 2015-12-22: Sum; 2016-01-12: Average; 2016-01-18: Variance
| 9.1187 | 33.83 | 42.9487 | 2015-12-08: 2015-12-02; 2015-12-21: 2; 2015-12-22: 91.1505; 2016-01-12: 45.5752; 2016-01-18: 5.7386
2016-01-18 | 9.3399 | 34.15 | 43.4899 | 2015-12-08: 2015-12-08; 2015-12-21: 3; 2015-12-22: 120.5059; 2016-01-12: 40.1686; 2016-01-18: 3.7614
| 9.6076 | 34.33 | 43.9376 | 2015-12-08: 2015-12-21; 2015-12-21: 4; 2015-12-22: 153.6397; 2016-01-12: 38.4099; 2016-01-18: 2.5459
| 10.0747 | 34.47 | 44.5447 | 2015-12-08: 2015-12-22; 2015-12-21: 4; 2015-12-22: 170.553; 2016-01-12: 42.6383; 2016-01-18: 0.15174
| 9.8961 | 34.39 | 44.2861 | 2015-12-08: 2016-01-12; 2015-12-21: 4; 2015-12-22: 171.5066; 2016-01-12: 42.8767; 2016-01-18: 0.054626
2016-01-20 | 8.5052 | 33.74 | 42.2452 | 2015-12-08: 2016-01-18; 2015-12-21: 4; 2015-12-22: 176.2583; 2016-01-12: 44.0646; 2016-01-18: 0.20866
| 9.3468 | 35.35 | 44.6968 | 2015-12-08: 2016-01-20; 2015-12-21: 4; 2015-12-22: 169.6289; 2016-01-12: 42.4072; 2016-01-18: 2.631
| 9.7584 | 32.01 | 41.7684 | 2015-12-08: 2016-02-02; 2015-12-21: 4; 2015-12-22: 169.4679; 2016-01-12: 42.367; 2016-01-18: 0.36432
| 8.5485 | 32.37 | 40.9185 | 2015-12-08: 2016-02-09; 2015-12-21: 4; 2015-12-22: 169.8336; 2016-01-12: 42.4584; 2016-01-18: 1.6296
2016-02-02 | 10.1017 | 31.52 | 41.6217 | 2015-12-08: 2016-02-16; 2015-12-21: 4; 2015-12-22: 156.9186; 2016-01-12: 39.2297; 2016-01-18: 0.15997
| 9.6343 | 33.46 | 43.0943 | 2016-01-12: 42.0196
| 9.2114 | 33.1 | 42.3114
| 9.6606 | 32.78 | 42.4406 | 2015-12-08: ANOVA
2016-02-09 | 8.3357 | 32.9 | 41.2357 |  | Source of Variation | SS | df | MS | F | P-value | F crit
| 8.2148 | 34.71 | 42.9248 |  | Between Groups | 141.12 | 9 | 15.68 | 11.5992 | 3.3294e-07 | 2.2501
| 8.6144 | 33.02 | 41.6344 | 2015-12-08: Within Groups; 2015-12-21: 36.4989; 2015-12-22: 27; 2016-01-12: 1.3518
| 9.6286 | 34.41 | 44.0386
2016-02-16 | 6.6727 | 32.51 | 39.1827 | 2015-12-08: Total; 2015-12-21: 177.6189; 2015-12-22: 36
| 6.4852 | 33.27 | 39.7552 | 2016-01-12: mg/100 g
| 7.1791 | 32.02 | 39.1991 | 2015-12-08: Repeatabiltiy; 2015-12-22: Sr; 2016-01-12: 1.1627
| 7.1417 | 31.64 | 38.7817 | 2015-12-22: Sb; 2016-01-12: 1.8926
[5 empty] | Reproducibility |  | Si | 2.2212
NIST
Date | Bound Inositol Result(mg/100g) | Free Inositol Result(mg/100g) | Total Inositol Result(mg/100g)
2015-05-26 | 1 |  |  | 2015-12-08: NIST
| 0.77 |  |  | 2015-12-02: Total NIST
2015-05-28 | 0.76 |  |  | 2015-12-08 | 2015-12-21 | 2016-01-12 | 2016-01-18 | 2016-01-20 | 2016-02-02 | 2016-02-09 | 2016-02-16
2015-12-08 | 0.83232 | 41.57 | 42.4023 | 42.4023 | 37.8312 | 43.2619 | 40.6985 | 40.7454 | 38.3381 | 41.0275 | 40.0537
2015-12-21 | 0.91119 | 36.92 | 37.8312 |  |  | 42.9909 | 40.9598 | 41.7699 | 39.7653 | 42.8368 | 39.7404
2015-12-22 | 1.0113 | 31.56 | 32.5713
2015-01-12 | 0.69193 | 42.57 | 43.2619
| 0.75091 | 42.24 | 42.9909 | 2015-12-08: Anova: Single Factor
2015-01-18 | 0.92848 | 39.77 | 40.6985
| 0.94978 | 40.01 | 40.9598 | 2015-12-08: SUMMARY; 2016-01-18: mg/100 g
2015-01-20 | 0.88544 | 39.86 | 40.7454 | Groups | Count | Sum | Average | Variance
| 0.82994 | 40.94 | 41.7699 | 2015-12-08: 2015-12-08; 2015-12-21: 1; 2016-01-12: 42.4023; 2016-01-18: 42.4023
2015-02-02 | 0.79809 | 37.54 | 38.3381 | 2015-12-08: 2015-12-21; 2015-12-21: 1; 2016-01-12: 37.8312; 2016-01-18: 37.8312
| 0.71531 | 39.05 | 39.7653 | 2016-01-12 | 2 | 86.2528 | 43.1264 | 0.036726
2015-02-09 | 0.76751 | 40.26 | 41.0275 | 2016-01-18 | 2 | 81.6583 | 40.8291 | 0.034138
| 0.78685 | 42.05 | 42.8368 | 2016-01-20 | 2 | 82.5154 | 41.2577 | 0.52481
2015-02-16 | 0.45373 | 39.6 | 40.0537 | 2016-02-02 | 2 | 78.1034 | 39.0517 | 1.0185
| 0.53039 | 39.21 | 39.7404 | 2016-02-09 | 2 | 83.8644 | 41.9322 | 1.6369
|  |  |  | 2016-02-16 | 2 | 79.7941 | 39.8971 | 0.049089
ANOVA
|  |  |  | Source of Variation | SS | df | MS | F | P-value | F crit
|  |  |  | Between Groups | 32.8272 | 7 | 4.6896 | 8.5263 | 0.0092215 | 4.2067
|  |  |  | Within Groups | 3.3001 | 6 | 0.55001
|  |  |  | Total | 36.1273 | 13
mg/100 g
|  |  |  | Repeatabiltiy |  | Sr | 0.74163
[6 empty] | Sb | 1.0173
|  |  |  | Reproducibility |  | Si | 1.2589