## Precision & MU
n | 17
k | 2
Reproducibility sd | 0.93957 | Reproducibility from ANOVA
Mean | 63.6941 | Average of samples
Precision | 1.8791 | =k*Reproducibility | k=expansion coefficient
Bias | -0.77143 | Insignificant | Average of Bias*mean of sample/100
UoBias | 5.8192 |  | Uncertainty*mean/100
MU ± | 11.7892 | mg/kg |  | 63.694 | mg/kg
18.51%
Replicates | 2.6575
Rounding | 0.1
Horwitz MU | 63.68±10.9 | 17.12%
| 3 | 6.8065
| 2*U(Bias) < abs(averageBias) | significant
| 2*U(Bias) > abs(averageBias) | insignificant

## Repeatability
|  |  |  | n | 2
|  |  |  | k | 4.3027
|  |  |  | Repeatability sd | 0
|  |  |  | Mean | #DIV/0!
|  |  |  | Duplicates | 0
|  |  |  | Rounding | #NUM!
Matrix | Ref | REPORTED_UNITS | A | B | Diff | sqr | Omit | Outlier
[5 empty] | 0 | 0 | No | No
[5 empty] | 0 | 0 | No | Yes

## ANOVA
A | B | C | D | E | F | G
63.9 | 63.9 | 63.5 | 65 | 64.9 | 63.4 | 62.9
63.7 |  | 62.4 | 63.1 | 64.8 | 62.5 | 63.3
|  |  | 62.5 | 64.5 | 65.4 | 63.1
Anova: Single Factor
SUMMARY
Groups | Count | Sum | Average | Variance
A | 2 | 127.6 | 63.8 | 0.02 | 0.03141%
B | 1 | 63.9 | 63.9 |  | 0%
C | 2 | 125.9 | 62.95 | 0.605 | 0.95%
D | 3 | 190.6 | 63.5333 | 1.7033 | 2.675%
E | 3 | 194.2 | 64.7333 | 0.043333 | 0.06805%
F | 3 | 191.3 | 63.7667 | 2.2033 | 3.46%
G | 3 | 189.3 | 63.1 | 0.04 | 0.06281%
|  | Average | 63.6833 | mg/kg | 1.035%
ANOVA
Source of Variation | SS | df | MS | F | P-value | F crit
Between Groups | 5.5644 | 6 | 0.9274 | 1.0777 | 0.43588 | 3.2172
Within Groups | 8.605 | 10 | 0.8605
Total | 14.1694 | 16
Standard Deviation
Repeatability |  | Sr | 0.92763
|  | Sb | 0.14933
Reproducibility |  | Si | 0.93957
Intermediate precision |  |  | 1.8791 | mg/kg
2.951%
F < F critical | homogeneity and reproducibility
P-value > 0.05 | greater 95% confidence

## Bias
| U(Ref Conc), % | 0.62295
| Average Bias, % | -1.2111
| sd (Bias), % | 9.115
| U(Bias), % | 9.1362
| Bias Significance | Insignificant
| 2*U(Bias) < abs(aveBias) | significant
| 2*U(Bias) > abs(aveBias) | insignificant
Type | Reference | Date | Unit | Matrix | Reference Mean | sd | n | Lab Result | % sd | Bias | % Bias
CRM | NIST 1849a |  | mg/kg | IF | 68.2 | 1.9 | 20 | 77.61 | 2.7859 | 9.41 | 13.7977 | Batch# | Sample# | SRM 1849a NIST (µg/100g)
CRM | NIST 1849a |  | mg/kg | IF | 68.2 | 1.9 | 20 | 75.22 | 2.7859 | 7.02 | 10.2933 |  |  | FA | Accuracy | 5-Me THF | Accuracy
CRM | NIST 1849a |  | mg/kg | IF | 68.2 | 1.9 | 20 | 73.41 | 2.7859 | 7.02 | 10.2933 | 2014-12-13 | 1 | 263.6 | 115% | 10.1 | 120% | 273.7
CRM | NIST 1849a |  | mg/kg | IF | 68.2 | 1.9 | 20 | 76.46 | 2.7859 | 5.21 | 7.6393 |  | 2 | 233.9 | 102% | 9.2 | 109% | 243.1
CRM | NIST 1849a |  | mg/kg | IF | 68.2 | 1.9 | 20 | 73.63 | 2.7859 | 8.26 | 12.1114 | 2014-12-20 | 3 | 252.4 | 110% | 9.5 | 113% | 261.9
CRM | NIST 1849a |  | mg/kg | IF | 68.2 | 1.9 | 20 | 73.17 | 2.7859 | 5.43 | 7.9619 |  | 4 | 255.3 | 111% | 9.9 | 118% | 265.2
CRM | NIST 1849a |  | mg/kg | IF | 68.2 | 1.9 | 20 | 71.06 | 2.7859 | 4.97 | 7.2874 | 2015-04-01 | 5 | 256 | 112% | 9.7 | 116% | 265.7
CRM | NIST 1849a |  | mg/kg | IF | 68.2 | 1.9 | 20 | 73.17 | 2.7859 | 2.86 | 4.1935 |  | 6 | 233 | 102% | 10.2 | 121% | 243.2
CRM | NIST 1849a |  | mg/kg | IF | 68.2 | 1.9 | 20 | 64.7 | 2.7859 | 4.97 | 7.2874 |  | 7 | 211 | 92% | 10.5 | 125% | 221.5
CRM | NIST 1849a |  | mg/kg | IF | 68.2 | 1.9 | 20 | 66 | 2.7859 | -3.5 | -5.132 | 2015-04-20 | 8 | 221.3 | 97% | 9.4 | 112% | 230.7
CRM | NIST 1849a |  | mg/kg | IF | 68.2 | 1.9 | 20 | 62 | 2.7859 | -2.2 | -3.2258 |  | 9 | 206.9 | 90% | 11.9 | 142% | 218.8
CRM | NIST 1849a |  | mg/kg | IF | 68.2 | 1.9 | 20 | 61.4 | 2.7859 | -6.2 | -9.0909 |  | 10 | 216.6 | 95% | 9.7 | 115% | 226.3
CRM | NIST 1849a |  | mg/kg | IF | 68.2 | 1.9 | 20 | 62.7 | 2.7859 | -3.5 | -5.132 | 2015-07-01 | 11 | 210.2 | 92% | NA | NA | #VALUE!
CRM | NIST 1849a |  | mg/kg | IF | 68.2 | 1.9 | 20 | 54.9 | 2.7859 | -2.2 | -3.2258 |  | 12 | 224.5 | 98% | NA | NA | #VALUE!
CRM | NIST 1849a |  | mg/kg | IF | 68.2 | 1.9 | 20 | 66.6 | 2.7859 | -6.2 | -9.0909 | 2015-07-02 | 13 | 260.8 | 114% | 11.1 | 132% | 271.9
CRM | NIST 1849a |  | mg/kg | IF | 68.2 | 1.9 | 20 | 68.3 | 2.7859 | -6.8 | -9.9707 |  | 14 | 248.1 | 108% | 11.4 | 136% | 259.5
CRM | NIST 1849a |  | mg/kg | IF | 68.2 | 1.9 | 20 | 65.6 | 2.7859 | -5.5 | -8.0645 | 2015-08-05 | 15 | 249.5 | 109% | 11.4 | 136% | 260.9
CRM | NIST 1849a |  | mg/kg | IF | 68.2 | 1.9 | 20 | 66.7 | 2.7859 | -13.3 | -19.5015 |  | 16 | 245.5 | 107% | 11.8 | 140% | 257.3
CRM | NIST 1849a |  | mg/kg | IF | 68.2 | 1.9 | 20 | 67.3 | 2.7859 | -1.6 | -2.346 | 2015-08-10 | 17 | 219.8 | 96% | 10.9 | 130% | 230.7
CRM | NIST 1849a |  | mg/kg | IF | 68.2 | 1.9 | 20 | 58 | 2.7859 | 0.1 | 0.14663 |  | 18 | 221.8 | 97% | 10.9 | 130% | 232.7
CRM | NIST 1849a |  | mg/kg | IF | 68.2 | 1.9 | 20 | 58.5 | 2.7859 | -2.6 | -3.8123 | 2015-08-12 | 19 | 201.7 | 88% | 10.3 | 123% | 212
CRM | NIST 1849a |  | mg/kg | IF | 68.2 | 1.9 | 20 | 63.7 | 2.7859 | -1.5 | -2.1994 |  | 20 | 197.8 | 86% | 10 | 120% | 207.8
CRM | NIST 1849a |  | mg/kg | IF | 68.2 | 1.9 | 20 | 66.8 | 2.7859 | -0.9 | -1.3196 | 2015-08-14 | 21 | 226 | 99% | 10.2 | 122% | 236.2
CRM | NIST 1849a |  | mg/kg | IF | 68.2 | 1.9 | 20 | 65.9 | 2.7859 | -10.2 | -14.956 |  | 22 | 234.8 | 102% | 12.2 | 145% | 247
CRM | NIST 1849a |  | mg/kg | IF | 68.2 | 1.9 | 20 | 65.7 | 2.7859 | -9.7 | -14.2229 | 2015-08-18 | 23 | 214.9 | 94% | 10.6 | 127% | 225.5
[12 empty] | 2015-08-21 | 24 | 249.4 | 109% | 10 | 119% | 259.4

## Linearity
Conc | Response
0 | 0
1 | 1.02
2 | 2.03
4 | 4.1
6 | 6.1
10 | 10.05

## LOD-LOR
sd | 0.0040222 | #DIV/0! | #DIV/0! | #DIV/0!
LOD | 0.012066 | #DIV/0! | #DIV/0! | #DIV/0!
LOR | 0.040222 | #DIV/0! | #DIV/0! | #DIV/0!
Result | 1 | 2 | 3 | 4
1 | 0.01
2 | 0.02
3 | 0.01
4 | 0.015
5 | 0.014
6 | 0.01
7 | 0.011
8 | 0.02
9 | 0.014
10 | 0.018