## Precision & MU
n | 22
k | 2
Reproducibility sd | 9.6435 | Reproducibility from ANOVA
Mean | 133.2409 | Average of samples
Precision | 19.2871 | =k*Reproducibility | k=expansion coefficient
Bias | 1.1756 | Insignificant | Average of Bias*mean of sample/100
UoBias | 11.3587 |  | Uncertainty*mean/100 | 133.241 | ug/100 g
|  |  |  | 1332.41 | ug/kg
MU ± | 29.8005 | ug/100 g |  | 1.3324 | mg/kg
| 22.37% | Howitz MU | 1.332±0.408 | 30.65%
Replicates | 27.2761
Rounding | 1
| 3 | 17.2053
| 2*U(Bias) < abs(averageBias) | significant
| 2*U(Bias) > abs(averageBias) | insignificant

## Repeatability
|  |  |  | n | 2
|  |  |  | k | 4.3027
|  |  |  | Repeatability sd | 0
|  |  |  | Mean | #DIV/0!
|  |  |  | Duplicates | 0
|  |  |  | Rounding | #NUM!
Matrix | Ref | REPORTED_UNITS | A | B | Diff | sqr | Omit | Outlier
[5 empty] | 0 | 0 | No | No
[5 empty] | 0 | 0 | No | Yes

## ANOVA
2014-12-13 | 2014-12-20 | 2015-04-01 | 2015-04-20 | 2015-08-04 | 2015-08-05 | 2015-08-10 | 2015-08-12 | 2015-08-14 | 2015-08-18 | 2015-08-21 | Batch# | Sample | IRM001A (µg/100 g) |  |  | ug/100 g
130.9 | 136 | 138 | 130.2 | 146.8 | 144.8 | 130.6 | 124 | 127.9 | 125 | 141.5 |  |  | FA | 5-Me THF | Total | A | B | C | D | E | F | G | H | I | J | K
128.5 | 128.7 | 125 | 129.6 | 132.2 | 162 | 129.6 | 117.7 | 128.2 |  |  | 2014-12-13 | 1 | 98.1 | 32.8 | 130.9 | 130.9 | 136 | 138 | 130.2 | 146.8 | 144.8 | 130.6 | 124 | 127.9 | 125 | 141.5
2015-08-04: 139.9 |  | 2 | 100.5 | 28 | 128.5 | 128.5 | 128.7 | 125 | 129.6 | 132.2 | 162 | 129.6 | 117.7 | 128.2
2015-08-04: 134.2 | 2014-12-20 | 3 | 101.6 | 34.4 | 136 |  |  |  |  | 139.9
134.2
2014-12-13: Anova: Single Factor |  | 4 | 97.2 | 31.5 | 128.7
[11 empty] | 2015-04-01 | 5 | 100 | 38 | 138
2014-12-13: SUMMARY |  | 6 | 90 | 35 | 125
2014-12-13: Groups; 2014-12-20: Count; 2015-04-01: Sum; 2015-04-20: Average; 2015-08-04: Variance; 2015-08-05: CV | 2015-04-20 | 7 | 87.9 | 42.3 | 130.2
2014-12-13: 2014-12-13; 2014-12-20: 2; 2015-04-01: 259.4; 2015-04-20: 129.7; 2015-08-04: 2.88; 2015-08-05: 0.021689 |  | 8 | 88.6 | 41 | 129.6
2014-12-13: 2014-12-20; 2014-12-20: 2; 2015-04-01: 264.7; 2015-04-20: 132.35; 2015-08-04: 26.645; 2015-08-05: 0.20066 |  | 9 | 92.6 | 38.4 | 131
2014-12-13: 2015-04-01; 2014-12-20: 2; 2015-04-01: 263; 2015-04-20: 131.5; 2015-08-04: 84.5; 2015-08-05: 0.63637 | 2015-08-04 | 10 | 107.3 | 39.5 | 146.8
2014-12-13: 2015-04-20; 2014-12-20: 2; 2015-04-01: 259.8; 2015-04-20: 129.9; 2015-08-04: 0.18; 2015-08-05: 0.0013556 |  | 11 | 93.4 | 38.8 | 132.2
2014-12-13: 2015-08-04; 2014-12-20: 4; 2015-04-01: 553.1; 2015-04-20: 138.275; 2015-08-04: 42.9425; 2015-08-05: 0.3234 |  | 12 | 98.9 | 41 | 139.9
2014-12-13: 2015-08-05; 2014-12-20: 2; 2015-04-01: 306.8; 2015-04-20: 153.4; 2015-08-04: 147.92; 2015-08-05: 1.114 |  | 13 | 93.9 | 40.3 | 134.2
2014-12-13: 2015-08-10; 2014-12-20: 2; 2015-04-01: 260.2; 2015-04-20: 130.1; 2015-08-04: 0.5; 2015-08-05: 0.0037655 | 2015-08-05 | 14 | 106.7 | 38.1 | 144.8
2014-12-13: 2015-08-12; 2014-12-20: 2; 2015-04-01: 241.7; 2015-04-20: 120.85; 2015-08-04: 19.845; 2015-08-05: 0.14945 |  | 15 | 120 | 42 | 162
2014-12-13: 2015-08-14; 2014-12-20: 2; 2015-04-01: 256.1; 2015-04-20: 128.05; 2015-08-04: 0.045; 2015-08-05: 0.0003389 | 2015-08-10 | 16 | 94.2 | 36.4 | 130.6
2014-12-13: 2015-08-18; 2014-12-20: 1; 2015-04-01: 125; 2015-04-20: 125; 2015-08-05: 0 |  | 17 | 94.6 | 35 | 129.6
2014-12-13: 2015-08-21; 2014-12-20: 1; 2015-04-01: 141.5; 2015-04-20: 141.5; 2015-08-05: 0 | 2015-08-12 | 18 | 91.8 | 32.2 | 124
2015-04-01: Average; 2015-04-20: 132.7841; 2015-08-04: ug/100 g; 2015-08-05: 0.22282 |  | 19 | 84.7 | 33 | 117.7
[11 empty] | 2015-08-14 | 20 | 91.7 | 36.2 | 127.9
2014-12-13: ANOVA |  | 21 | 91 | 37.2 | 128.2
Source of Variation | SS | df | MS | F | P-value | F crit |  |  |  |  | 2015-08-18 | 22 | 89.6 | 35.4 | 125
Between Groups | 1486.0107 | 10 | 148.6011 | 3.9738 | 0.016384 | 2.8536 |  |  |  |  | 2015-08-21 | 23 | 108.3 | 33.2 | 141.5
Within Groups | 411.3425 | 11 | 37.3948
Total | 1897.3532 | 21
sd
Repeatability |  | Sr | 6.1151
|  | Sb | 7.4568
Reproducibility |  | Si | 9.6435
Intermediate precision |  |  | 19.2871 | ug/100 g
15.43%
F < F critical | homogeneity and reproducibility
P-value > 0.05 | greater 95% confidence

## Bias
| U(Ref Conc), % | 1.7158
| Average Bias, % | 0.88232
| sd (Bias), % | 8.3505
| U(Bias), % | 8.5249
| Bias Significance | Insignificant
| 2*U(Bias) < abs(aveBias) | significant
| 2*U(Bias) > abs(aveBias) | insignificant
Type | Reference | Date | Unit | Matrix | Reference Mean | sd | n | Lab Result | % sd | Bias | % Bias
CRM | NIST 1849a |  | ug/100g | IF | 241.1 | 18.5 | 20 | 273.7 | 7.6732 | 32.6 | 13.5214 | Batch# | Sample# | SRM 1849a NIST (µg/100g)
CRM | NIST 1849a |  | ug/100g | IF | 241.1 | 18.5 | 20 | 243.1 | 7.6732 | 2 | 0.82953 |  |  | FA | Accuracy | 5-Me THF | Accuracy
CRM | NIST 1849a |  | ug/100g | IF | 241.1 | 18.5 | 20 | 261.9 | 7.6732 | 20.8 | 8.6271 | 2014-12-13 | 1 | 263.6 | 115% | 10.1 | 120% | 273.7
CRM | NIST 1849a |  | ug/100g | IF | 241.1 | 18.5 | 20 | 265.2 | 7.6732 | 24.1 | 9.9959 |  | 2 | 233.9 | 102% | 9.2 | 109% | 243.1
CRM | NIST 1849a |  | ug/100g | IF | 241.1 | 18.5 | 20 | 265.7 | 7.6732 | 24.6 | 10.2032 | 2014-12-20 | 3 | 252.4 | 110% | 9.5 | 113% | 261.9
CRM | NIST 1849a |  | ug/100g | IF | 241.1 | 18.5 | 20 | 243.2 | 7.6732 | 2.1 | 0.87101 |  | 4 | 255.3 | 111% | 9.9 | 118% | 265.2
CRM | NIST 1849a |  | ug/100g | IF | 241.1 | 18.5 | 20 | 221.5 | 7.6732 | -19.6 | -8.1294 | 2015-04-01 | 5 | 256 | 112% | 9.7 | 116% | 265.7
CRM | NIST 1849a |  | ug/100g | IF | 241.1 | 18.5 | 20 | 230.7 | 7.6732 | -10.4 | -4.3136 |  | 6 | 233 | 102% | 10.2 | 121% | 243.2
CRM | NIST 1849a |  | ug/100g | IF | 241.1 | 18.5 | 20 | 218.8 | 7.6732 | -22.3 | -9.2493 |  | 7 | 211 | 92% | 10.5 | 125% | 221.5
CRM | NIST 1849a |  | ug/100g | IF | 241.1 | 18.5 | 20 | 226.3 | 7.6732 | -14.8 | -6.1385 | 2015-04-20 | 8 | 221.3 | 97% | 9.4 | 112% | 230.7
CRM | NIST 1849a |  | ug/100g | IF | 241.1 | 18.5 | 20 | 271.9 | 7.6732 | 30.8 | 12.7748 |  | 9 | 206.9 | 90% | 11.9 | 142% | 218.8
CRM | NIST 1849a |  | ug/100g | IF | 241.1 | 18.5 | 20 | 259.5 | 7.6732 | 18.4 | 7.6317 |  | 10 | 216.6 | 95% | 9.7 | 115% | 226.3
CRM | NIST 1849a |  | ug/100g | IF | 241.1 | 18.5 | 20 | 260.9 | 7.6732 | 19.8 | 8.2124 | 2015-07-01 | 11 | 210.2 | 92% | NA | NA | #VALUE!
CRM | NIST 1849a |  | ug/100g | IF | 241.1 | 18.5 | 20 | 257.3 | 7.6732 | 16.2 | 6.7192 |  | 12 | 224.5 | 98% | NA | NA | #VALUE!
CRM | NIST 1849a |  | ug/100g | IF | 241.1 | 18.5 | 20 | 230.7 | 7.6732 | -10.4 | -4.3136 | 2015-07-02 | 13 | 260.8 | 114% | 11.1 | 132% | 271.9
CRM | NIST 1849a |  | ug/100g | IF | 241.1 | 18.5 | 20 | 232.7 | 7.6732 | -8.4 | -3.484 |  | 14 | 248.1 | 108% | 11.4 | 136% | 259.5
CRM | NIST 1849a |  | ug/100g | IF | 241.1 | 18.5 | 20 | 212 | 7.6732 | -29.1 | -12.0697 | 2015-08-05 | 15 | 249.5 | 109% | 11.4 | 136% | 260.9
CRM | NIST 1849a |  | ug/100g | IF | 241.1 | 18.5 | 20 | 207.8 | 7.6732 | -33.3 | -13.8117 |  | 16 | 245.5 | 107% | 11.8 | 140% | 257.3
CRM | NIST 1849a |  | ug/100g | IF | 241.1 | 18.5 | 20 | 236.2 | 7.6732 | -4.9 | -2.0324 | 2015-08-10 | 17 | 219.8 | 96% | 10.9 | 130% | 230.7
CRM | NIST 1849a |  | ug/100g | IF | 241.1 | 18.5 | 20 | 247 | 7.6732 | 5.9 | 2.4471 |  | 18 | 221.8 | 97% | 10.9 | 130% | 232.7
CRM | NIST 1849a |  | ug/100g | IF | 241.1 | 18.5 | 20 | 225.5 | 7.6732 | -15.6 | -6.4703 | 2015-08-12 | 19 | 201.7 | 88% | 10.3 | 123% | 212
CRM | NIST 1849a |  | ug/100g | IF | 241.1 | 18.5 | 20 | 259.4 | 7.6732 | 18.3 | 7.5902 |  | 20 | 197.8 | 86% | 10 | 120% | 207.8
[12 empty] | 2015-08-14 | 21 | 226 | 99% | 10.2 | 122% | 236.2
[13 empty] | 22 | 234.8 | 102% | 12.2 | 145% | 247
[12 empty] | 2015-08-18 | 23 | 214.9 | 94% | 10.6 | 127% | 225.5
[12 empty] | 2015-08-21 | 24 | 249.4 | 109% | 10 | 119% | 259.4

## Linearity
Conc | Response
0 | 0
1 | 1.02
2 | 2.03
4 | 4.1
6 | 6.1
10 | 10.05

## LOD-LOR
sd | 0.0040222 | #DIV/0! | #DIV/0! | #DIV/0!
LOD | 0.012066 | #DIV/0! | #DIV/0! | #DIV/0!
LOR | 0.040222 | #DIV/0! | #DIV/0! | #DIV/0!
Result | 1 | 2 | 3 | 4
1 | 0.01
2 | 0.02
3 | 0.01
4 | 0.015
5 | 0.014
6 | 0.01
7 | 0.011
8 | 0.02
9 | 0.014
10 | 0.018

## Latest_1922016
IRM001A-1
Date | Bound Inositol Result(mg/100g) | Free Inositol Result(mg/100g) | Total Inositol Result(mg/100g)
2015-12-02 | 10.4913 | 33.39 | 43.8813
| 9.1591 | 38.11 | 47.2691
9.4486
| 9.3861 |  |  |  | IRM001A-1
2015-12-08 | 8.3694 | 31.17 | 39.5394
| 7.8145 | 34.53 | 42.3445 | 2015-12-02 | 2015-12-08 | 2015-12-21 | 2015-12-22 | 2016-01-12 | 2016-01-18 | 2016-01-20 | 2016-02-02 | 2016-02-09 | 2016-02-16
| 8.4619 | 30.16 | 38.6219 | 43.8813 | 39.5394 | 36.2376 | 42.5642 | 42.8258 | 43.4899 | 42.2452 | 41.6217 | 41.2357 | 39.1827
2015-12-21 | 7.5776 | 28.66 | 36.2376 | 47.2691 | 42.3445 | 39.5779 | 42.8939 | 43.1453 | 43.9376 | 44.6968 | 43.0943 | 42.9248 | 39.7552
| 8.0179 | 31.56 | 39.5779 |  | 38.6219 | 39.6368 | 42.9764 | 42.5869 | 44.5447 | 41.7684 | 42.3114 | 41.6344 | 39.1991
| 7.4968 | 32.14 | 39.6368 |  |  | 38.1873 | 42.1186 | 42.9487 | 44.2861 | 40.9185 | 42.4406 | 44.0386 | 38.7817
| 6.4373 | 31.75 | 38.1873
2015-12-22 | 11.4042 | 31.16 | 42.5642
| 11.7239 | 31.17 | 42.8939
| 11.1964 | 31.78 | 42.9764
| 11.4986 | 30.62 | 42.1186 | 2015-12-08: Anova: Single Factor
2016-01-12 | 8.5758 | 34.25 | 42.8258
| 8.8153 | 34.33 | 43.1453 | 2015-12-08: SUMMARY; 2016-01-12: mg/100 g
| 8.5969 | 33.99 | 42.5869 | 2015-12-08: Groups; 2015-12-21: Count; 2015-12-22: Sum; 2016-01-12: Average; 2016-01-18: Variance
| 9.1187 | 33.83 | 42.9487 | 2015-12-08: 2015-12-02; 2015-12-21: 2; 2015-12-22: 91.1505; 2016-01-12: 45.5752; 2016-01-18: 5.7386
2016-01-18 | 9.3399 | 34.15 | 43.4899 | 2015-12-08: 2015-12-08; 2015-12-21: 3; 2015-12-22: 120.5059; 2016-01-12: 40.1686; 2016-01-18: 3.7614
| 9.6076 | 34.33 | 43.9376 | 2015-12-08: 2015-12-21; 2015-12-21: 4; 2015-12-22: 153.6397; 2016-01-12: 38.4099; 2016-01-18: 2.5459
| 10.0747 | 34.47 | 44.5447 | 2015-12-08: 2015-12-22; 2015-12-21: 4; 2015-12-22: 170.553; 2016-01-12: 42.6383; 2016-01-18: 0.15174
| 9.8961 | 34.39 | 44.2861 | 2015-12-08: 2016-01-12; 2015-12-21: 4; 2015-12-22: 171.5066; 2016-01-12: 42.8767; 2016-01-18: 0.054626
2016-01-20 | 8.5052 | 33.74 | 42.2452 | 2015-12-08: 2016-01-18; 2015-12-21: 4; 2015-12-22: 176.2583; 2016-01-12: 44.0646; 2016-01-18: 0.20866
| 9.3468 | 35.35 | 44.6968 | 2015-12-08: 2016-01-20; 2015-12-21: 4; 2015-12-22: 169.6289; 2016-01-12: 42.4072; 2016-01-18: 2.631
| 9.7584 | 32.01 | 41.7684 | 2015-12-08: 2016-02-02; 2015-12-21: 4; 2015-12-22: 169.4679; 2016-01-12: 42.367; 2016-01-18: 0.36432
| 8.5485 | 32.37 | 40.9185 | 2015-12-08: 2016-02-09; 2015-12-21: 4; 2015-12-22: 169.8336; 2016-01-12: 42.4584; 2016-01-18: 1.6296
2016-02-02 | 10.1017 | 31.52 | 41.6217 | 2015-12-08: 2016-02-16; 2015-12-21: 4; 2015-12-22: 156.9186; 2016-01-12: 39.2297; 2016-01-18: 0.15997
| 9.6343 | 33.46 | 43.0943 | 2016-01-12: 42.0196
| 9.2114 | 33.1 | 42.3114
| 9.6606 | 32.78 | 42.4406 | 2015-12-08: ANOVA
2016-02-09 | 8.3357 | 32.9 | 41.2357 |  | Source of Variation | SS | df | MS | F | P-value | F crit
| 8.2148 | 34.71 | 42.9248 |  | Between Groups | 141.12 | 9 | 15.68 | 11.5992 | 3.3294e-07 | 2.2501
| 8.6144 | 33.02 | 41.6344 | 2015-12-08: Within Groups; 2015-12-21: 36.4989; 2015-12-22: 27; 2016-01-12: 1.3518
| 9.6286 | 34.41 | 44.0386
2016-02-16 | 6.6727 | 32.51 | 39.1827 | 2015-12-08: Total; 2015-12-21: 177.6189; 2015-12-22: 36
| 6.4852 | 33.27 | 39.7552 | 2016-01-12: mg/100 g
| 7.1791 | 32.02 | 39.1991 | 2015-12-08: Repeatabiltiy; 2015-12-22: Sr; 2016-01-12: 1.1627
| 7.1417 | 31.64 | 38.7817 | 2015-12-22: Sb; 2016-01-12: 1.8926
[5 empty] | Reproducibility |  | Si | 2.2212
NIST
Date | Bound Inositol Result(mg/100g) | Free Inositol Result(mg/100g) | Total Inositol Result(mg/100g)
2015-05-26 | 1 |  |  | 2015-12-08: NIST
| 0.77 |  |  | 2015-12-02: Total NIST
2015-05-28 | 0.76 |  |  | 2015-12-08 | 2015-12-21 | 2016-01-12 | 2016-01-18 | 2016-01-20 | 2016-02-02 | 2016-02-09 | 2016-02-16
2015-12-08 | 0.83232 | 41.57 | 42.4023 | 42.4023 | 37.8312 | 43.2619 | 40.6985 | 40.7454 | 38.3381 | 41.0275 | 40.0537
2015-12-21 | 0.91119 | 36.92 | 37.8312 |  |  | 42.9909 | 40.9598 | 41.7699 | 39.7653 | 42.8368 | 39.7404
2015-12-22 | 1.0113 | 31.56 | 32.5713
2015-01-12 | 0.69193 | 42.57 | 43.2619
| 0.75091 | 42.24 | 42.9909 | 2015-12-08: Anova: Single Factor
2015-01-18 | 0.92848 | 39.77 | 40.6985
| 0.94978 | 40.01 | 40.9598 | 2015-12-08: SUMMARY; 2016-01-18: mg/100 g
2015-01-20 | 0.88544 | 39.86 | 40.7454 | Groups | Count | Sum | Average | Variance
| 0.82994 | 40.94 | 41.7699 | 2015-12-08: 2015-12-08; 2015-12-21: 1; 2016-01-12: 42.4023; 2016-01-18: 42.4023
2015-02-02 | 0.79809 | 37.54 | 38.3381 | 2015-12-08: 2015-12-21; 2015-12-21: 1; 2016-01-12: 37.8312; 2016-01-18: 37.8312
| 0.71531 | 39.05 | 39.7653 | 2016-01-12 | 2 | 86.2528 | 43.1264 | 0.036726
2015-02-09 | 0.76751 | 40.26 | 41.0275 | 2016-01-18 | 2 | 81.6583 | 40.8291 | 0.034138
| 0.78685 | 42.05 | 42.8368 | 2016-01-20 | 2 | 82.5154 | 41.2577 | 0.52481
2015-02-16 | 0.45373 | 39.6 | 40.0537 | 2016-02-02 | 2 | 78.1034 | 39.0517 | 1.0185
| 0.53039 | 39.21 | 39.7404 | 2016-02-09 | 2 | 83.8644 | 41.9322 | 1.6369
|  |  |  | 2016-02-16 | 2 | 79.7941 | 39.8971 | 0.049089
ANOVA
|  |  |  | Source of Variation | SS | df | MS | F | P-value | F crit
|  |  |  | Between Groups | 32.8272 | 7 | 4.6896 | 8.5263 | 0.0092215 | 4.2067
|  |  |  | Within Groups | 3.3001 | 6 | 0.55001
|  |  |  | Total | 36.1273 | 13
mg/100 g
|  |  |  | Repeatabiltiy |  | Sr | 0.74163
[6 empty] | Sb | 1.0173
|  |  |  | Reproducibility |  | Si | 1.2589